## Tabelle1
Document | Medical device file according to ISO 13485:2016 | Ref.: ISO 13485:2016 Chapter 4.2.3 Section | Technical documentation according to MDR/IVDR | Ref.: MDR (respective annex and section) | Ref.: IVDR (respective annex and section) | FDA: DMR, DHF, DHR
General description of the device | X | a) | X (incl. names) | II.1.1.a) | II.1.1. a)
Variants | X*) |  | X | II.1.1. i) | II.1.1. l)
Language versions | X*) |  | X | II.2. | II.2.
Intended accessories and combinations with other devices | X*) |  | X | II.1.1. h) | II.1.1. m)
Intended purpose and intended use | X | a) | X (incl. user, patient group, (contra-) indications) | II.1.1 a), c) | II.1.1 a), c) | 21 CFR part 820.30
Warnings |  |  | X | II.1.1. c) | n.a.
Qualification as medical device |  |  | X | II.1.1.e) | II.1.1. e)
Classification | X*) |  | X | II.1.1. f) | II.1.1. f)
Mode of action, scientifically demonstrated if necessary, functionality (IVD: measuring principle, mode of operation) |  |  | X | II.1.1. d) | II.1.1. d)
Description of novel features |  |  | X | II.1.1 g) | n.a.
Reference to previous and similar devices |  |  | X | II.1.2. a), b) | II.1.2. a), b) | e.g. for 510(k) approval
Description of the development stages |  |  | X | II.3. a) | II.3.1. | 21 CFR part 820.30 b)
Device specifications, e.g., drawings, calculations, components | X | b) | X (incl. dimensions, performance attributes) | II.1.1. j), l) | II.1.1. g)-k), II.3.1. | 21 CFR part 820.181 a)
Software specifications | X*) |  | X | II.6.1 b) | II.3.1.b)-d) u. II.6.4 | 21 CFR part 820.181 a)
Relevant GSPR |  |  | X | II.4. a) | II.4. a)
Materials used, design drawings | X*) |  | X (incl. those in contact with the body) | II.1.1 j) | II.1.1. g), II.3.1. a)-c) | 21 CFR part 820.181 a)
Medicinal products used (IVD: in relation to companion diagnostics (CDx)) | X*) |  | X | II.1.1. k), II.6.2. a) | II.1.1. c)ix) (for CDx)
Tissues of human or animal origin used |  |  | XX | II.6.2 b) | II.6.5. b)
Toxicity, biological safety, and biocompatibility (IVD: Analysis performance incl. sample type) | X*) |  | X | II.6.1 b), II.6.2. d) | II.6.1. | (21 CFR part 820.181 c)), 21 CFR part 820.184 d)
Physical, chemical, microbiological, electrical safety |  |  | X | II.6.1 b) | II.4. | (21 CFR part 820.181 c)), 21 CFR part 820.184 d)
Absorption, metabolism, excretion, interaction, tolerance of substances in the body |  |  | X | II.6.2 c) | n.a.
Lifetime, stability, shelf life |  |  | X | II.6.1 b) | II.6.3. | 21 CFR part 820.184 d)
Measurement accuracy (in case of measuring function) |  |  | X | II.6.2. f) | II.6.5. c) | 21 CFR part 820.184 d)
Interoperability, compatibility |  |  | X | II.6.2.g) | II.6.5. d) | (21 CFR part 820.181 c))
Software V&V results |  |  | X | II.6.1 b) | II.6.4. | 21 CFR part 820.30
Specifications for manufacture, e.g., requirements for facilities, infrastructure, production methods | X | c) | (incl. auxiliary materials, processes, validation) | II.3. b) | II.3.2. | 21 CFR part 820.181 b)
Quality assurance requirements, e.g., procedures, equipment, acceptance criteria | X*) |  | X | II.6.1 (a.o.) | II.6. | 21 CFR part 820.181 c)
Details of suppliers and contractors with their activities |  |  | X | II.3. c) | II.3.2. b)
Packaging specifications, e.g., methods, processes, materials | X | c) | X | II.6.2. e) | II.3.2. a), II.6.5. a) | 21 CFR part 820.181 d)
Specifications (procedures) for measuring and monitoring | X | d) | X | II.3. b), II.6.1. d), III | III | (21 CFR part 820.181 c))
Requirements for post-market surveillance (PMS), particular the PMS Plan |  |  | X | III | III
Design outputs | X*) |  | X | II.6. (a.o.) | II.6. (u.a.)
Tests, e.g., laboratory, simulation, animal testing |  |  | X | II.6.1 a) | II.6. | (21 CFR part 820.181a)), 21 CFR part 820.184 d)
Clinical data (IVD: Performance evaluation) | X*) |  | X | II.6.1 c) | II.6.1., 6.2.
Clinical evaluation, incl. CER and CEP (IVD: Clinical performance and clinical evidence, PER) |  |  | X | II.6.1 c) | II.6.2.
PMCF plan (IVD: PMPF plan) |  |  | X | II.6.1 d) | III.1. b)
Risk management, e.g., risk analysis, risk-minimizing measures, list of residual risks, risk-benefit evaluation | X*) |  | X | II.5. a), b) | II.5.
Storage requirements | X | c) |  |  | II.6.3.1., II.6.3.2.
Specifications for transport | X*) | c) |  |  | II.6.3.3.
Specifications for distribution | X | c)
Requirements for installation | X | e) |  |  |  | 21 CFR part 820.181e)
Requirements for maintenance and servicing | X | f) |  |  |  | 21 CFR part 820.181d)
Instructions for use | X | c) | X | II.2. | II.2. b)
Other labeling | X | c) | X (packing) | II.2. | II.2. a)
UDI | X*) |  | X (Basic UDI) | II 1.1. b) | II 1.1. b)
Changes to the device during its lifetime and related V&V | X*) |  | X | Article 10 | Article 10
List of standards (or “other methods”) used to demonstrate conformity | X*) |  | X | II.4. b), c) | II.4. b), c)
(Documents to) demonstrate conformity with GSPR | X*) |  | X | II.4. d) | II.4. d)